## DemoTest
name | enabled | main_menu_label | main_menu_xy | side_menu_label | side_menu_xy | data_view_label | data_view_xy | fixed_text | meas_low | meas_high | meas_unit | ratio_low | ratio_high | ratio_text | ratio_unit | reset | timestamp_count | modbus_addr | modbus_type | modbus_low | modbus_high | modbus_unit | modbus_aggre_selection | notes
Test case name | TRUE/FALSE/init | ConfigTouch member hint | Example: home; 120,50 | Side menu hint | Side menu xy | Data/toggle hint | Data/toggle xy | Expected fixed text, comma separated | Measurement lower bounds, semicolon separated | Measurement upper bounds, semicolon separated | Measurement unit | Ratio lower bounds, semicolon separated | Ratio upper bounds, semicolon separated | Ratio expected text, comma separated | Ratio unit | TRUE runs max/min reset before case | count; margin seconds, blank skips | Modbus Map 문서 주소 입력. 내부에서 -1 적용. 세미콜론 다중 | float/uint16/int16/uint32/int32. 1개면 모든 addr에 적용 | 통신값 하한. 세미콜론 다중 | 통신값 상한. 세미콜론 다중 | 통신값 단위 표시용 | A7300 전용. A2700은 blank | Source notes
초기화 | init
VOLTAGE-RMS-LL | TRUE | touch_main_menu_1 | meter; home; 100, 85 | touch_side_menu_1 | 80, 135 | touch_toggle_ll | 410, 150 | RMS Voltage,L-L,L-N,Min,Max,AB,BC,CA,Average | 189.0; 189.0; 189.0; 189.0 | 191.0; 191.0; 191.0; 191.0 | V | [6 empty] | 15011; 15013; 15015; 15017 | float | 189.0; 189.0; 189.0; 189.0 | 191.0; 191.0; 191.0; 191.0 | V | 1 | v1.0.1 demo_process.test_steps[1]; variant=base; Modbus addr_* not used by current xlsx loader
VOLTAGE-RMS-LL-Min | TRUE |  |  |  |  | touch_toggle_min | 620, 150 | RMS Voltage,L-L,L-N,Min,Max,AB,BC,CA,Average | 189.0; 189.0; 189.0; 189.0 | 191.0; 191.0; 191.0; 191.0 | V |  |  |  |  | TRUE | 4; 20 | 31211; 31213; 31215; 31217 | float | 189.0; 189.0; 189.0; 189.0 | 191.0; 191.0; 191.0; 191.0 | V | 255 | v1.0.1 demo_process.test_steps[1]; variant=min; Modbus addr_* not used by current xlsx loader
VOLTAGE-RMS-LL-Max | TRUE |  |  |  |  | touch_toggle_max | 720, 150 | RMS Voltage,L-L,L-N,Min,Max,AB,BC,CA,Average | 189.0; 189.0; 189.0; 189.0 | 191.0; 191.0; 191.0; 191.0 | V |  |  |  |  | TRUE | 4; 20 | 31011; 31013; 31015; 31017 | float | 189.0; 189.0; 189.0; 189.0 | 191.0; 191.0; 191.0; 191.0 | V | 255 | v1.0.1 demo_process.test_steps[1]; variant=max; Modbus addr_* not used by current xlsx loader
VOLTAGE-RMS-LN | TRUE |  |  |  |  | touch_toggle_ln | 510, 150 | RMS Voltage,L-L,L-N,Min,Max,A,B,C,Average | 109.0; 109.0; 109.0; 109.0 | 111.0; 111.0; 111.0; 111.0 | V | [6 empty] | 15001; 15003; 15005; 15007 | float | 109.0; 109.0; 109.0; 109.0 | 111.0; 111.0; 111.0; 111.0 | V | 1 | v1.0.1 demo_process.test_steps[2]; variant=base; Modbus addr_* not used by current xlsx loader
VOLTAGE-RMS-LN-Min | TRUE |  |  |  |  | touch_toggle_min | 620, 150 | RMS Voltage,L-L,L-N,Min,Max,A,B,C,Average | 109.0; 109.0; 109.0; 109.0 | 111.0; 111.0; 111.0; 111.0 | V |  |  |  |  | TRUE | 4; 20 | 31201; 31203; 31205; 31207 | float | 109.0; 109.0; 109.0; 109.0 | 111.0; 111.0; 111.0; 111.0 | V | 255 | v1.0.1 demo_process.test_steps[2]; variant=min; Modbus addr_* not used by current xlsx loader
VOLTAGE-RMS-LN-Max | TRUE |  |  |  |  | touch_toggle_max | 720, 150 | RMS Voltage,L-L,L-N,Min,Max,A,B,C,Average | 109.0; 109.0; 109.0; 109.0 | 111.0; 111.0; 111.0; 111.0 | V |  |  |  |  | TRUE | 4; 20 | 31001; 31003; 31005; 31007 | float | 109.0; 109.0; 109.0; 109.0 | 111.0; 111.0; 111.0; 111.0 | V | 255 | v1.0.1 demo_process.test_steps[2]; variant=max; Modbus addr_* not used by current xlsx loader
VOLTAGE-Fundamental-LL | TRUE |  |  | touch_side_menu_2 | 80, 180 |  |  | Fund. Volt.,L-L,L-N,Min,Max,AB,BC,CA,Average | 189.0; 189.0; 189.0; 189.0 | 191.0; 191.0; 191.0; 191.0 | V | [6 empty] | 15039; 15041; 15043; 15045 | float | 189.0; 189.0; 189.0; 189.0 | 191.0; 191.0; 191.0; 191.0 | V | 1 | v1.0.1 demo_process.test_steps[3]; variant=base; Modbus addr_* not used by current xlsx loader
VOLTAGE-Fundamental-LL-Min | TRUE |  |  |  |  | touch_toggle_min | 620, 150 | Fund. Volt.,L-L,L-N,Min,Max,AB,BC,CA,Average | 189.0; 189.0; 189.0; 189.0 | 191.0; 191.0; 191.0; 191.0 | V |  |  |  |  | TRUE | 4; 20 | 31239; 31241; 31243; 31245 | float | 189.0; 189.0; 189.0; 189.0 | 191.0; 191.0; 191.0; 191.0 | V | 255 | v1.0.1 demo_process.test_steps[3]; variant=min; Modbus addr_* not used by current xlsx loader
VOLTAGE-Fundamental-LL-Max | TRUE |  |  |  |  | touch_toggle_max | 720, 150 | Fund. Volt.,L-L,L-N,Min,Max,AB,BC,CA,Average | 189.0; 189.0; 189.0; 189.0 | 191.0; 191.0; 191.0; 191.0 | V |  |  |  |  | TRUE | 4; 20 | 31039; 31041; 31043; 31045 | float | 189.0; 189.0; 189.0; 189.0 | 191.0; 191.0; 191.0; 191.0 | V | 255 | v1.0.1 demo_process.test_steps[3]; variant=max; Modbus addr_* not used by current xlsx loader
VOLTAGE-Fundamental-LN | TRUE |  |  |  |  | touch_toggle_ln | 510, 150 | Fund. Volt.,L-L,L-N,Min,Max,A,B,C,Average | 109.0; 109.0; 109.0; 109.0 | 111.0; 111.0; 111.0; 111.0 | V | [6 empty] | 15029; 15031; 15033; 15035 | float | 109.0; 109.0; 109.0; 109.0 | 111.0; 111.0; 111.0; 111.0 | V | 1 | v1.0.1 demo_process.test_steps[5]; variant=base; Modbus addr_* not used by current xlsx loader
VOLTAGE-Fundamental-LN-Min | TRUE |  |  |  |  | touch_toggle_min | 620, 150 | Fund. Volt.,L-L,L-N,Min,Max,A,B,C,Average | 109.0; 109.0; 109.0; 109.0 | 111.0; 111.0; 111.0; 111.0 | V |  |  |  |  | TRUE | 4; 20 | 31229; 31231; 31233; 31235 | float | 109.0; 109.0; 109.0; 109.0 | 111.0; 111.0; 111.0; 111.0 | V | 255 | v1.0.1 demo_process.test_steps[5]; variant=min; Modbus addr_* not used by current xlsx loader
VOLTAGE-Fundamental-LN-Max | TRUE |  |  |  |  | touch_toggle_max | 720, 150 | Fund. Volt.,L-L,L-N,Min,Max,A,B,C,Average | 109.0; 109.0; 109.0; 109.0 | 111.0; 111.0; 111.0; 111.0 | V |  |  |  |  | TRUE | 4; 20 | 31029; 31031; 31033; 31035 | float | 109.0; 109.0; 109.0; 109.0 | 111.0; 111.0; 111.0; 111.0 | V | 255 | v1.0.1 demo_process.test_steps[5]; variant=max; Modbus addr_* not used by current xlsx loader
VOLTAGE-THD-LL | TRUE |  |  | touch_side_menu_3 | 80, 225 |  |  | Total Harmonic Distortion,L-L,L-N,Max,AB,BC,CA | 2.0; 2.0; 2.0 | 4.0; 4.0; 4.0 | % | [6 empty] | 15065; 15067; 15069 | float | 2.0; 2.0; 2.0 | 4.0; 4.0; 4.0 | % | 1 | v1.0.1 demo_process.test_steps[6]; variant=base; Modbus addr_* not used by current xlsx loader
VOLTAGE-THD-LL-Max | TRUE |  |  |  |  | touch_toggle_max | 720, 150 | Total Harmonic Distortion,L-L,L-N,Max,AB,BC,CA | 2.0; 2.0; 2.0 | 4.0; 4.0; 4.0 | % |  |  |  |  | TRUE | 3; 20 | 31105; 31107; 31109 | float | 2.0; 2.0; 2.0 | 4.0; 4.0; 4.0 | % | 255 | v1.0.1 demo_process.test_steps[6]; variant=max; Modbus addr_* not used by current xlsx loader
VOLTAGE-THD-LN | TRUE |  |  |  |  | touch_toggle_thd_ln; touch_toggle_max | 620, 150; 720, 150 | Total Harmonic Distortion,L-L,L-N,Max,A,B,C | 2.0; 2.0; 2.0 | 4.0; 4.0; 4.0 | % | [6 empty] | 15059; 15061; 15063 | float | 2.0; 2.0; 2.0 | 4.0; 4.0; 4.0 | % | 1 | v1.0.1 demo_process.test_steps[7]; variant=base; Modbus addr_* not used by current xlsx loader
VOLTAGE-THD-LN-Max | TRUE |  |  |  |  | touch_toggle_max | 720, 150 | Total Harmonic Distortion,L-L,L-N,Max,A,B,C | 2.0; 2.0; 2.0 | 4.0; 4.0; 4.0 | % |  |  |  |  | TRUE | 3; 20 | 31099; 31101; 31103 | float | 2.0; 2.0; 2.0 | 4.0; 4.0; 4.0 | % | 255 | v1.0.1 demo_process.test_steps[7]; variant=max; Modbus addr_* not used by current xlsx loader
VOLTAGE-Frequency | TRUE |  |  | touch_side_menu_4 | 80, 270 |  |  | Frequency,Min,Max,Frequency | 59.8 | 60.1 | Hz | [6 empty] | 15057 | float | 59.8 | 60.1 | Hz | 1 | v1.0.1 demo_process.test_steps[8]; variant=base; Modbus addr_* not used by current xlsx loader
VOLTAGE-Frequency-Min | TRUE |  |  |  |  | touch_toggle_min | 620, 150 | Frequency,Min,Max,Frequency | 59.8 | 60.1 | Hz |  |  |  |  | TRUE | 1; 20 | 31297 | float | 59.8 | 60.1 | Hz | 255 | v1.0.1 demo_process.test_steps[8]; variant=min; Modbus addr_* not used by current xlsx loader
VOLTAGE-Frequency-Max | TRUE |  |  |  |  | touch_toggle_max | 720, 150 | Frequency,Min,Max,Frequency | 59.8 | 60.1 | Hz |  |  |  |  | TRUE | 1; 20 | 31097 | float | 59.8 | 60.1 | Hz | 255 | v1.0.1 demo_process.test_steps[8]; variant=max; Modbus addr_* not used by current xlsx loader
VOLTAGE-Residual | TRUE |  |  | touch_side_menu_5 | 80, 315 |  |  | Residual Voltage,Min,Max,RMS,Fund. | 6.0; 0.7 | 8.0; 2.0 | V | [6 empty] | 15009; 15037 | float | 6.0; 0.7 | 8.0; 2.0 | V | 1 | v1.0.1 demo_process.test_steps[9]; variant=base; Modbus addr_* not used by current xlsx loader
VOLTAGE-Residual-Min | TRUE |  |  |  |  | touch_toggle_min | 620, 150 | Residual Voltage,Min,Max,RMS,Fund. | 6.0; 0.7 | 8.0; 2.0 | V |  |  |  |  | TRUE | 2; 20 | 31209; 31237 | float | 6.0; 0.7 | 8.0; 2.0 | V | 255 | v1.0.1 demo_process.test_steps[9]; variant=min; Modbus addr_* not used by current xlsx loader
VOLTAGE-Residual-Max | TRUE |  |  |  |  | touch_toggle_max | 720, 150 | Residual Voltage,Min,Max,RMS,Fund. | 6.0; 0.7 | 8.0; 2.0 | V |  |  |  |  | TRUE | 2; 20 | 31009; 31037 | float | 6.0; 0.7 | 8.0; 2.0 | V | 255 | v1.0.1 demo_process.test_steps[9]; variant=max; Modbus addr_* not used by current xlsx loader
CURRENT-RMS | TRUE | touch_main_menu_2 | home; 260, 85 | touch_side_menu_1 | 80, 135 |  |  | RMS Current,Min,Max,A,B,C,Average | 24.0; 24.0; 24.0; 24.0 | 26.0; 26.0; 26.0; 26.0 | A | 49.0; 49.0; 49.0; 49.0 | 51.0; 51.0; 51.0; 51.0 |  | % |  |  | 15019; 15021; 15023; 15025 | float | 24.0; 24.0; 24.0; 24.0 | 26.0; 26.0; 26.0; 26.0 | A | 1 | v1.0.1 demo_process.test_steps[10]; variant=base; Modbus addr_* not used by current xlsx loader
CURRENT-RMS-Min | TRUE |  |  |  |  | touch_toggle_min | 620, 150 | RMS Current,Min,Max,A,B,C,Average | 24.0; 24.0; 24.0; 24.0 | 26.0; 26.0; 26.0; 26.0 | A | 49.0; 49.0; 49.0; 49.0 | 51.0; 51.0; 51.0; 51.0 |  | % | TRUE | 4; 20 | 31219; 31221; 31223; 31225 | float | 24.0; 24.0; 24.0; 24.0 | 26.0; 26.0; 26.0; 26.0 | A | 255 | v1.0.1 demo_process.test_steps[10]; variant=min; Modbus addr_* not used by current xlsx loader
CURRENT-RMS-Max | TRUE |  |  |  |  | touch_toggle_max | 720, 150 | RMS Current,Min,Max,A,B,C,Average | 24.0; 24.0; 24.0; 24.0 | 26.0; 26.0; 26.0; 26.0 | A | 49.0; 49.0; 49.0; 49.0 | 51.0; 51.0; 51.0; 51.0 |  | % | TRUE | 4; 20 | 31019; 31021; 31023; 31025 | float | 24.0; 24.0; 24.0; 24.0 | 26.0; 26.0; 26.0; 26.0 | A | 255 | v1.0.1 demo_process.test_steps[10]; variant=max; Modbus addr_* not used by current xlsx loader
CURRENT-Fundamental | TRUE |  |  | touch_side_menu_2 | 80, 180 |  |  | Fundamental Current,Min,Max,A,B,C,Average | 24.0; 24.0; 24.0; 24.0 | 26.0; 26.0; 26.0; 26.0 | A | 49.0; 49.0; 49.0; 49.0 | 51.0; 51.0; 51.0; 51.0 |  | % |  |  | 15047; 15049; 15051; 15053 | float | 24.0; 24.0; 24.0; 24.0 | 26.0; 26.0; 26.0; 26.0 | A | 1 | v1.0.1 demo_process.test_steps[11]; variant=base; Modbus addr_* not used by current xlsx loader
CURRENT-Fundamental-Min | TRUE |  |  |  |  | touch_toggle_min | 620, 150 | Fundamental Current,Min,Max,A,B,C,Average | 24.0; 24.0; 24.0; 24.0 | 26.0; 26.0; 26.0; 26.0 | A | 49.0; 49.0; 49.0; 49.0 | 51.0; 51.0; 51.0; 51.0 |  | % | TRUE | 4; 20 | 31247; 31249; 31251; 31253 | float | 24.0; 24.0; 24.0; 24.0 | 26.0; 26.0; 26.0; 26.0 | A | 255 | v1.0.1 demo_process.test_steps[11]; variant=min; Modbus addr_* not used by current xlsx loader
CURRENT-Fundamental-Max | TRUE |  |  |  |  | touch_toggle_max | 720, 150 | Fundamental Current,Min,Max,A,B,C,Average | 24.0; 24.0; 24.0; 24.0 | 26.0; 26.0; 26.0; 26.0 | A | 49.0; 49.0; 49.0; 49.0 | 51.0; 51.0; 51.0; 51.0 |  | % | TRUE | 4; 20 | 31047; 31049; 31051; 31053 | float | 24.0; 24.0; 24.0; 24.0 | 26.0; 26.0; 26.0; 26.0 | A | 255 | v1.0.1 demo_process.test_steps[11]; variant=max; Modbus addr_* not used by current xlsx loader
CURRENT-Demand | TRUE |  |  | touch_side_menu_3 | 80, 225 |  |  | Demand Current,Peak,A,B,C,Average | 24.0; 24.0; 24.0; 24.0 | 26.0; 26.0; 26.0; 26.0 | A | 49.0; 49.0; 49.0; 49.0 | 51.0; 51.0; 51.0; 51.0 |  | % |  |  | 15171; 15173; 15175; 15177 | float | 24.0; 24.0; 24.0; 24.0 | 26.0; 26.0; 26.0; 26.0 | A | 1 | v1.0.1 demo_process.test_steps[12]; variant=base; Modbus addr_* not used by current xlsx loader
CURRENT-Demand-Max | TRUE |  |  |  |  | touch_toggle_max | 720, 150 | Demand Current,Peak,A,B,C,Average | 24.0; 24.0; 24.0; 24.0 | 26.0; 26.0; 26.0; 26.0 | A | 49.0; 49.0; 49.0; 49.0 | 51.0; 51.0; 51.0; 51.0 |  | % | TRUE | 4; 20 | 31167; 31169; 31171; 31173 | float | 24.0; 24.0; 24.0; 24.0 | 26.0; 26.0; 26.0; 26.0 | A | 255 | v1.0.1 demo_process.test_steps[12]; variant=max; Modbus addr_* not used by current xlsx loader
CURRENT-THD | TRUE |  |  | touch_side_menu_4 | 80, 270 |  |  | Total Harmonic Distortion,Max,A,B,C | 1.5; 1.5; 1.5 | 2.5; 2.5; 2.5 | % | [6 empty] | 15071; 15073; 15075 | float | 1.5; 1.5; 1.5 | 2.5; 2.5; 2.5 | % | 1 | v1.0.1 demo_process.test_steps[13]; variant=base; Modbus addr_* not used by current xlsx loader
CURRENT-THD-Max | TRUE |  |  |  |  | touch_toggle_max | 720, 150 | Total Harmonic Distortion,Max,A,B,C | 1.5; 1.5; 1.5 | 2.5; 2.5; 2.5 | % |  |  |  |  | TRUE | 3; 20 | 31111; 31113; 31115 | float | 1.5; 1.5; 1.5 | 2.5; 2.5; 2.5 | % | 255 | v1.0.1 demo_process.test_steps[13]; variant=max; Modbus addr_* not used by current xlsx loader
CURRENT-TDD | TRUE |  |  | touch_side_menu_5 | 80, 315 |  |  | Total Demand Distortion,Max,A,B,C | 0.5; 0.5; 0.5 | 2.5; 2.5; 2.5 | % | [6 empty] | 15077; 15079; 15081 | float | 0.5; 0.5; 0.5 | 2.5; 2.5; 2.5 | % | 1 | v1.0.1 demo_process.test_steps[14]; variant=base; Modbus addr_* not used by current xlsx loader
CURRENT-TDD-Max | TRUE |  |  |  |  | touch_toggle_max | 720, 150 | Total Demand Distortion,Max,A,B,C | 0.5; 0.5; 0.5 | 2.5; 2.5; 2.5 | % |  |  |  |  | TRUE | 3; 20 | 31117; 31119; 31121 | float | 0.5; 0.5; 0.5 | 2.5; 2.5; 2.5 | % | 255 | v1.0.1 demo_process.test_steps[14]; variant=max; Modbus addr_* not used by current xlsx loader
CURRENT-Crest Factor | TRUE |  |  | touch_side_menu_6 | 80, 360 |  |  | Crest Factor,Max,A,B,C | 1.4; 1.4; 1.4 | 1.5; 1.5; 1.5 | [7 empty] | 15083; 15085; 15087 | float | 1.4; 1.4; 1.4 | 1.5; 1.5; 1.5 |  | 1 | v1.0.1 demo_process.test_steps[15]; variant=base; Modbus addr_* not used by current xlsx loader
CURRENT-Crest Factor-Max | TRUE |  |  |  |  | touch_toggle_max | 720, 150 | Crest Factor,Max,A,B,C | 1.4; 1.4; 1.4 | 1.5; 1.5; 1.5 | [5 empty] | TRUE | 3; 20 | 31123; 31125; 31127 | float | 1.4; 1.4; 1.4 | 1.5; 1.5; 1.5 |  | 255 | v1.0.1 demo_process.test_steps[15]; variant=max; Modbus addr_* not used by current xlsx loader
CURRENT-K Factor | TRUE |  |  | touch_side_menu_7 | 80, 405 |  |  | K-Factor,Max,A,B,C | 1.1; 1.1; 1.1 | 1.5; 1.5; 1.5 | [7 empty] | 15089; 15091; 15093 | float | 1.1; 1.1; 1.1 | 1.5; 1.5; 1.5 |  | 1 | v1.0.1 demo_process.test_steps[16]; variant=base; Modbus addr_* not used by current xlsx loader
CURRENT-K Factor-Max | TRUE |  |  |  |  | touch_toggle_max | 720, 150 | K-Factor,Max,A,B,C | 1.1; 1.1; 1.1 | 1.5; 1.5; 1.5 | [5 empty] | TRUE | 3; 20 | 31129; 31131; 31133 | float | 1.1; 1.1; 1.1 | 1.5; 1.5; 1.5 |  | 255 | v1.0.1 demo_process.test_steps[16]; variant=max; Modbus addr_* not used by current xlsx loader
CURRENT-Residual | TRUE |  |  | touch_side_menu_8 | 80, 450 |  |  | Residual Current,Min,Max,RMS,Fund. | 0.5; 0.1 | 1.0; 0.5 | A | [6 empty] | 15027; 15055 | float | 0.5; 0.1 | 1.0; 0.5 | A | 1 | v1.0.1 demo_process.test_steps[17]; variant=base; Modbus addr_* not used by current xlsx loader
CURRENT-Residual-Min | TRUE |  |  |  |  | touch_toggle_min | 620, 150 | Residual Current,Min,Max,RMS,Fund. | 0.5; 0.1 | 1.0; 0.5 | A |  |  |  |  | TRUE | 2; 20 | 31227; 31255 | float | 0.5; 0.1 | 1.0; 0.5 | A | 255 | v1.0.1 demo_process.test_steps[17]; variant=min; Modbus addr_* not used by current xlsx loader
CURRENT-Residual-Max | TRUE |  |  |  |  | touch_toggle_max | 720, 150 | Residual Current,Min,Max,RMS,Fund. | 0.5; 0.1 | 1.0; 0.5 | A |  |  |  |  | TRUE | 2; 20 | 31027; 31055 | float | 0.5; 0.1 | 1.0; 0.5 | A | 255 | v1.0.1 demo_process.test_steps[17]; variant=max; Modbus addr_* not used by current xlsx loader
POWER-Active | TRUE | touch_main_menu_3 | home; 390, 85 | touch_side_menu_1 | 80, 135 |  |  | Active Power,Min,Max,A,B,C,Total | 2.3; 2.3; 2.3; 6.9 | 2.4; 2.4; 2.4; 7.2 | kW | 40.0; 40.0; 40.0; 40.0 | 45.0; 45.0; 45.0; 45.0 |  | % |  |  | 15095; 15097; 15099; 15101 | float | 2.3; 2.3; 2.3; 6.9 | 2.4; 2.4; 2.4; 7.2 | kW | 1 | v1.0.1 demo_process.test_steps[18]; variant=base; Modbus addr_* not used by current xlsx loader
POWER-Active-Min | TRUE |  |  |  |  | touch_toggle_min | 620, 150 | Active Power,Min,Max,A,B,C,Total | 2.3; 2.3; 2.3; 6.9 | 2.4; 2.4; 2.4; 7.2 | kW | 40.0; 40.0; 40.0; 40.0 | 45.0; 45.0; 45.0; 45.0 |  | % | TRUE | 4; 20 | 31257; 31259; 31261; 31263 | float | 2.3; 2.3; 2.3; 6.9 | 2.4; 2.4; 2.4; 7.2 | kW | 255 | v1.0.1 demo_process.test_steps[18]; variant=min; Modbus addr_* not used by current xlsx loader
POWER-Active-Max | TRUE |  |  |  |  | touch_toggle_max | 720, 150 | Active Power,Min,Max,A,B,C,Total | 2.3; 2.3; 2.3; 6.9 | 2.4; 2.4; 2.4; 7.2 | kW | 40.0; 40.0; 40.0; 40.0 | 45.0; 45.0; 45.0; 45.0 |  | % | TRUE | 4; 20 | 31057; 31059; 31061; 31063 | float | 2.3; 2.3; 2.3; 6.9 | 2.4; 2.4; 2.4; 7.2 | kW | 255 | v1.0.1 demo_process.test_steps[18]; variant=max; Modbus addr_* not used by current xlsx loader
POWER-Reactive | TRUE |  |  | touch_side_menu_2 | 80, 180 |  |  | Reactive Power,Min,Max,A,B,C,Total | 1.3; 1.3; 1.3; 3.9 | 1.4; 1.4; 1.4; 4.2 | kVAR | 20.0; 20.0; 20.0; 20.0 | 30.0; 30.0; 30.0; 30.0 |  | % |  |  | 15103; 15105; 15107; 15109 | float | 1.3; 1.3; 1.3; 3.9 | 1.4; 1.4; 1.4; 4.2 | kVAR | 1 | v1.0.1 demo_process.test_steps[19]; variant=base; Modbus addr_* not used by current xlsx loader
POWER-Reactive-Min | TRUE |  |  |  |  | touch_toggle_min | 620, 150 | Reactive Power,Min,Max,A,B,C,Total | 1.3; 1.3; 1.3; 3.9 | 1.4; 1.4; 1.4; 4.2 | kVAR | 20.0; 20.0; 20.0; 20.0 | 30.0; 30.0; 30.0; 30.0 |  | % | TRUE | 4; 20 | 31265; 31267; 31269; 31271 | float | 1.3; 1.3; 1.3; 3.9 | 1.4; 1.4; 1.4; 4.2 | kVAR | 255 | v1.0.1 demo_process.test_steps[19]; variant=min; Modbus addr_* not used by current xlsx loader
POWER-Reactive-Max | TRUE |  |  |  |  | touch_toggle_max | 720, 150 | Reactive Power,Min,Max,A,B,C,Total | 1.3; 1.3; 1.3; 3.9 | 1.4; 1.4; 1.4; 4.2 | kVAR | 20.0; 20.0; 20.0; 20.0 | 30.0; 30.0; 30.0; 30.0 |  | % | TRUE | 4; 20 | 31065; 31067; 31069; 31071 | float | 1.3; 1.3; 1.3; 3.9 | 1.4; 1.4; 1.4; 4.2 | kVAR | 255 | v1.0.1 demo_process.test_steps[19]; variant=max; Modbus addr_* not used by current xlsx loader
POWER-Apparent | TRUE |  |  | touch_side_menu_3 | 80, 225 |  |  | Apparent Power,Min,Max,A,B,C,Total | 2.5; 2.5; 2.5; 7.5 | 2.9; 2.9; 2.9; 8.7 | kVA | 49.0; 49.0; 49.0; 49.0 | 51.0; 51.0; 51.0; 51.0 |  | % |  |  | 15111; 15113; 15115; 15117 | float | 2.5; 2.5; 2.5; 7.5 | 2.9; 2.9; 2.9; 8.7 | kVA | 1 | v1.0.1 demo_process.test_steps[20]; variant=base; Modbus addr_* not used by current xlsx loader
POWER-Apparent-Min | TRUE |  |  |  |  | touch_toggle_min | 620, 150 | Apparent Power,Min,Max,A,B,C,Total | 2.5; 2.5; 2.5; 7.5 | 2.9; 2.9; 2.9; 8.7 | kVA | 49.0; 49.0; 49.0; 49.0 | 51.0; 51.0; 51.0; 51.0 |  | % | TRUE | 4; 20 | 31273; 31275; 31277; 31279 | float | 2.5; 2.5; 2.5; 7.5 | 2.9; 2.9; 2.9; 8.7 | kVA | 255 | v1.0.1 demo_process.test_steps[20]; variant=min; Modbus addr_* not used by current xlsx loader
POWER-Apparent-Max | TRUE |  |  |  |  | touch_toggle_max | 720, 150 | Apparent Power,Min,Max,A,B,C,Total | 2.5; 2.5; 2.5; 7.5 | 2.9; 2.9; 2.9; 8.7 | kVA | 49.0; 49.0; 49.0; 49.0 | 51.0; 51.0; 51.0; 51.0 |  | % | TRUE | 4; 20 | 31073; 31075; 31077; 31079 | float | 2.5; 2.5; 2.5; 7.5 | 2.9; 2.9; 2.9; 8.7 | kVA | 255 | v1.0.1 demo_process.test_steps[20]; variant=max; Modbus addr_* not used by current xlsx loader
POWER-PF | TRUE |  |  | touch_side_menu_4 | 80, 270 |  |  | Power Factor,Min,Max,A,B,C,Total | 0.85; 0.85; 0.85; 0.85 | 0.9; 0.9; 0.9; 0.9 |  |  |  | Lag,Lag,Lag,Lag |  |  |  | 15119; 15121; 15123; 15125 | float | 0.85; 0.85; 0.85; 0.85 | 0.9; 0.9; 0.9; 0.9 |  | 1 | v1.0.1 demo_process.test_steps[21]; variant=base; Modbus addr_* not used by current xlsx loader
POWER-PF-Min | TRUE |  |  |  |  | touch_toggle_min | 620, 150 | Power Factor,Min,Max,A,B,C,Total | 0.85; 0.85; 0.85; 0.85 | 0.9; 0.9; 0.9; 0.9 |  |  |  | Lag,Lag,Lag,Lag |  | TRUE | 4; 20 | 31281; 31283; 31285; 31287 | float | 0.85; 0.85; 0.85; 0.85 | 0.9; 0.9; 0.9; 0.9 |  | 255 | v1.0.1 demo_process.test_steps[21]; variant=min; Modbus addr_* not used by current xlsx loader
POWER-PF-Max | TRUE |  |  |  |  | touch_toggle_max | 720, 150 | Power Factor,Min,Max,A,B,C,Total | 0.85; 0.85; 0.85; 0.85 | 0.9; 0.9; 0.9; 0.9 |  |  |  | Lag,Lag,Lag,Lag |  | TRUE | 4; 20 | 31081; 31083; 31085; 31087 | float | 0.85; 0.85; 0.85; 0.85 | 0.9; 0.9; 0.9; 0.9 |  | 255 | v1.0.1 demo_process.test_steps[21]; variant=max; Modbus addr_* not used by current xlsx loader
POWER-Demand | TRUE |  |  | touch_side_menu_5 | 80, 315 |  |  | Demand Active Power,Peak,A,B,C,Total | 2.3; 2.3; 2.3; 6.9 | 2.4; 2.4; 2.4; 7.2 | kW | 40.0; 40.0; 40.0; 40.0 | 45.0; 45.0; 45.0; 45.0 |  | % |  |  | 15179; 15181; 15183; 15185 | float | 2.3; 2.3; 2.3; 6.9 | 2.4; 2.4; 2.4; 7.2 | kW | 1 | v1.0.1 demo_process.test_steps[22]; variant=base; Modbus addr_* not used by current xlsx loader
POWER-Demand-Max | TRUE |  |  |  |  | touch_toggle_max | 720, 150 | Demand Active Power,Peak,A,B,C,Total | 2.3; 2.3; 2.3; 6.9 | 2.4; 2.4; 2.4; 7.2 | kW | 40.0; 40.0; 40.0; 40.0 | 45.0; 45.0; 45.0; 45.0 |  | % | TRUE | 4; 20 | 31175; 31177; 31179; 31181 | float | 2.3; 2.3; 2.3; 6.9 | 2.4; 2.4; 2.4; 7.2 | kW | 255 | v1.0.1 demo_process.test_steps[22]; variant=max; Modbus addr_* not used by current xlsx loader